## Slide 1: The Tandem heart system, device description and clinical results
By Robert G. Svitek
Summerized by Mahta Allsop

## Slide 2: What is the TandemHeart?
The TandemHeart System (CardiacAssist, Inc. Pittsburgh,
Pennsylvania) is an extracorporeal circulatory assist device
that drains blood from the left atrium (LA) and pumps it
back into the femoral artery to bypass the left ventricle.
It has four major components:
an extracorporeal centrifugal pump
a 21-Fr transseptal venous cannula (inflow)
a femoral arterial cannula (outflow)
a microprocessor-based controller.
The System can provide up to 5 L/min of blood flow
percutaneously and is CE marked for 30 days and approved by the
Food and Drug Administration (FDA) for use up to 6 hours

## Slide 4: How does the pump work?
a centrifugal blood pump that has a priming volume of 7 mL, weighs 280 g, and contains a hydrodynamic bearing to support the spinning impeller.
a motor chamber and a blood chamber that are separated by a polymeric seal.
motor is a brushless direct current motor consisting of a laminated stack of stainless steel plates wound with copper wire embedded in a heat-dissipative epoxy.
The rotor is centered in the stator, is spun by the motor through an electromagnetic coupling, and is attached directly to the impeller. Saline flows at 10 mL/h through the motor chamber and completely surrounds the rotor.

## Slide 5: How does the pump work?
Lift pads in the lower housing of the motor chamber provide thrust forces to stabilize the rotor in the axial direction.
The saline flows between a journal and the rotor toward the impeller to provide radial stability.
The impeller shaft passes through the center of the seal, and the saline (heparinized with 90,000 U/L) flows around the impeller shaft-seal interface to flush the area to prevent thrombus formation. The impeller contains six blades and rotates between 3000 and 7500 rpm (revolutions per minute) to provide flow rates from 1 to 8 L/min, depending on the size of the cannulae.

## Slide 6: Pump’s components
blood
impeller
seal
heparinized Saline solution
stator
rotor
Fluid bearing
journal

## Slide 7: The TandemHeart Escort is a microprocessor-based controller that powers the pump and has built-in alarms and diagnostics such as flow measurement that help assess pump function.

## Slide 8: Implantation procedure
The TandemHeart system can be inserted percutaneously in the catheterization laboratory through standard insertion techniques.
The femoral vein is accessed for transseptal puncture,
which is performed under fluoroscopic guidance using the Brockenbrough needle and a Mullins sheath.
Once the puncture is made in the fossa ovalis, heparin is given to achieve a target activated clotting time (ACT) of more than 300 seconds before inserting the cannula.

## Slide 9: Implantation procedure
After confirming the position of the Mullins sheath in the LA, a stiff wire is placed into the LA and the Mullins sheath is exchanged for the TandemHeart 14-Fr/21-Fr two-stage dilator. The dilator increases the septum puncture site diameter and enables the 21-Fr transseptal cannula to be inserted into the LA.
The cannula tip contains three radiopaque marker disks that can be viewed under fluoroscopy to confirm that the position of the cannula tip is in the LA. Transesophageal echocardiography has also been used to help with the insertion and cannula tip placement. The end of the cannula outside of the patient is sutured to the thigh to help prevent cannula dislodgement. Color Doppler has been used to visualize flow through the holes of the cannula tip to verify correct placement.

## Slide 10: Implantation procedure
The arterial cannula is placed in the femoral artery with the distal end sitting above the aortic bifurcation and is also sutured to the thigh to prevent cannula migration or dislodgement.
Once the transseptal and arterial cannulae are in place, the pump and cannulae are de-aired and connected via a wet-to-wet connection to remove all air from the circuit.
The pump is then connected to the TandemHeart controller, and its speed is adjusted to provide the desired level of support.

## Slide 11: Conclustion
The TandemHeart System has demonstrated the feasibility
of a left atrial to femoral artery bypass circuit. Clinical trials
in Europe and the United States provided evidence that the
TandemHeart can increase CI, MAP, and decrease PCWP which may be beneficial for patients with depressed cardiac function.
Future improvements to the cannula and pump may reduce
complications such as bleeding and increase the utility of the
device.

## Slide 12: Which Cardiac assist systems have I scanned?
Heartmate II
IABP (intra-aortic balloon pump)
ECMO (Extracorporeal membrane oxygenation)
Impella device